## Slide 1: Emissivity of black plated open honeycomb and black coatings at cryogenic temperatures
Makiko ANDO, Kosuke TANAKA, Keisuke SHINOZAKI, Hiroyuki SUGITA
(Japan Aerospace Exploration Agency)
Ikuma NISHIKAWA
(University of Tsukuba)

## Slide 2: Contents
Background and motivation
Approaches for emissivity enhancement
Emissivity measurements
Conclusions

## Slide 3: Background and motivation

## Slide 4: Background and motivation (1/2)
Ultra-sensitive astronomical observation missions
Optical sensors are required to be kept at cryogenic temperatures for thermal noise reduction.
Cooling system
Cryocoolers + Radiative cooling
Large-scale and long-term observation
Running out of cryogen
End of Mission
High emissivity surface at cryogenic temperatures is required.
Huge cryogen tank
Expendable cryogen cooling
No cryogen cooling
Cryogenic radiator surfaces of spacecraft
Deep space simulation at ground test facility (thermal vacuum test)
SPICA[1]
Lite BIRD
[1] H Ogawa, et al., ICES-2016-305 (2016)
©JAXA/ISAS

## Slide 5: Background and motivation (2/2)
Difficulty in achieving high emissivity at cryogenic temperatures
Emissivity of candidate surfaces was experimentally evaluated at 20-80K in this study.
Several kinds of coatings
Three-dimensional structure with/without coating
Dominant wavelengths of infrared radiation differ depending on temperature.
High emissivity surface at room temperature
High emissivity surface
at cryogenic temperature
≠
5-100mm at 300K
50-1000mm at 20K
5-100mm at 300K
50-1000mm at 20K
Dominant wavelengths of black-body radiation at various temperatures.
Longer and wider wavelength range should be considered.
e.g.)

## Slide 6: Approaches for emissivity enhancement

## Slide 7: Approach 1: Coatings
Emissivity enhancement by coating
Select a proper coating material and thickness
Candidate coatings in this study
Coating | Description
Black paint | Commonly used at room temperature to achieve high emissivity. Thickness: Usually on the order of several tens of mm.
Carbon nanotubes (CNT) | Extremely low reflectance at the wavelength range from UV to mid IR. Thickness: From submillimeter to millimeter.
Black plating | Common blackening method. Thickness: From several microns to several tens microns No organic compound Advantage of low outgassing

## Slide 8: Approach 2: Three-dimensional structures (1/2)
Cavity effect
Cavity resonance
Wavelength < Cavity aperture size
Cavity effect can be achieved.
No dependence on wavelength
repeatedly reflected and trapped inside the cavity.
Increase of
apparent emissivity
Incident light
Cavity
Cavity
. . .
Light at resonant wavelength is absorbed.
Increase of emissivity
at specific wavelength
Resonant wavelength(l)
2R = n x l/2
(n=1,2,…..)
2R

## Slide 9: Approach 2: Three-dimensional structures (2/2)
Emissivity enhancement by cylindrical cavity with diffuse material surface
(estimated by thermal analysis).
Cavity effect
Cylindrical cavity
L
R
Reach a plateau
Candidate structure in this study
Open honeycomb
Emissivity increase rate
Depending on surface emissivity
Increase rate reaches a plateau at an aspect ratio > 6.

## Slide 10: Emissivity Measurement

## Slide 11: Test pieces (1/2)
Name | Structure | Coating | 
 |  | Category | Thickness
Desothane®-Flat | Flat surface | Black paint (PPG Aerospace, Desothane®) | 130mm
CNT100-Flat |  | Carbon nanotubes | 100mm
CNT500-Flat |  | Carbon nanotubes | 500mm
TB-Flat |  | Black plating (Ebina Denka Kogyo, Tough Black®) | 20mm
TB-OH | Open honeycomb Cell size: 1/8” Cell height: 10mm | Black plating (Ebina Denka Kogyo, Tough Black®) | 20mm
AL-OH |  | No coating | -
1/8” (=3.2mm) could almost covers the dominant wavelength range of IR radiation down to 7K.
1/8” x 10mmH Aspect ratio: 6.3 (> 6)

## Slide 12: Test pieces (2/2)
Desothane®-Flat
CNT500-Flat
TB-Flat
TB-OH
AL-OH
Flat surface
Open honeycomb
Baseplate (Al or Ti)
50mm x 50mm x 3mmt
Coating
Baseplate (Al)
50mm x 50mm x 3mmt
Honeycomb core
10mm
A honeycomb core is adhered to the baseplate with STYCAST®.

## Slide 13: Measurement configuration
Test apparatus
Sample details (example of open honeycomb)
GHP (Guard heater plate)
Test sample
Measurement using calorimetric method
Low emissivity tape
Ultimate vacuum:
1.8×10-6Pa

## Slide 14: Defining equations
Test configuration and heat flow
Radiation heat transfer at steady state
where,
Symbols
A : area, m2
Q : heat transfer rate, W
T : temperature, K
e : emissivity, -
s : Stefan-Boltzmann constant, W/m2K4
Subscripts
eff : effective
H : heater
leak : leak
rad : radiation
ref : reference
s : sample
sh : shield
Approximation
Heat leak (measured in the heat leak test)
Major error source 1
Major error source 2
(1)
(2)
(3)
(4)
(5)

## Slide 15: Defining equations – Major error source (1/2)
Error due to the approximation.
Error due to the approximation
As/Ash = 1/200 in this study
Error due to the approximation: not always negligible.
Error varies depending on the shield emissivity, esh.
Emissivity of the 2nd shield, esh
Coated with black paint (Desothane®, 130mmt)
esh at 4K was not precisely identified.
Expected minimum value is 0.1 based on extrapolation using the emissivity measurement results above 20K.
Error due to the approximation could be +4.4% at maximum.
Symbols
A : area, m2
e : emissivity, -
Subscripts
eff : effective
s : sample
sh : shield

## Slide 16: Defining equations – Major error source (2/2)
Test configuration
Heat leak (measured in the heat leak test)
Conduction heat transfer by residual gas
Via Kevlar® threads and lead wires
Qleak is determined according to a preliminary heat leak test.
Same configuration and temperatures as for emissivity measurements.
Low emissivity reference sample is used (eref < 0.025).
The value of eref should be in the range from 0 to 0.025, but not identified at various cryogenic temperatures.
The uncertainty of eref should be considered in the error of es: +0.025 at maximum.
Qrad,ref
QH,ref
Ref. sample
Tref, Aref, eref
Symbols
A : area, m2 Q : heat transfer rate, W
T : temperature, K e : emissivity, -
s : Stefan-Boltzmann constant, W/m2K4
Subscripts
H : heater leak : leak
rad : radiation ref : reference
sh : sheild

## Slide 17: Result and discussion 1: Flat surface with coatings
Sample temperature [K]
Sample emissivity [-]
The emissivity of flat surface
Decrease significantly with temperature.
CNT500-Flat
Highest emissivity due to the large thickness of CNT coating.
Emissivity decreased at temperatures below 50K.
Decrease with temperature
Tough Black® had relatively high emissivity for its thickness of only 20mm.
Influence of wavelengths longer than the coating thickness.
Due to the complex texture at its surface.
©Ebina Denka Kogyo
Emissivity measurement results (flat surface with coatings).

## Slide 18: Result and discussion 2: Open honeycomb
Sample temperature [K]
Sample emissivity [-]
The emissivity of open honeycomb
Less dependent on temperature.
AL-OH (w/o black plating)
Emissivity enhanced up to 0.5 even when the material surface emissivity is very low.
TB-OH (w/ black plating)
Emissivity increased drastically by adding black plating.
Emissivity enhancing effect:
Specular reflection > Diffuse reflection
w/ black plating
w/o black plating
Emissivity measurement results (open honeycomb).
Open honeycomb + black plating
Promising solution to achieving high emissivity!

## Slide 19: Conclusions
The emissivity was measured using calorimetric methods at 20-80K.
Emissivity of flat surface with several kinds of black coatings
Strong dependence on temperature.
- Decreased significantly with temperature.
Emissivity of open honeycomb w/ or w/o black plating
Less dependent on temperature.
Enhanced up to 0.5 even without black coating, and 0.9 with black plating.
The combination of an open honeycomb and black plating
Promising solution to achieving high emissivity.

## Slide 20: Thank you for your kind attention!

## Slide 21: Back up slides

## Slide 22: Emissivity enhancing effect - Diffuse or Specular
M Kiuchi, Thermal Calculation Handbook, Japan Management Association, Published in 1988 (in Japanese)
Aspect Ratio (L/R) [-]
0
2
4
6
8
10
1.0
0.8
0.6
0.4
0.2
0.1
esurface=0.1
esurface=0.2
esurface=0.3
Apparent emissivity [-]
esurface=0.5
esurface=0.7
esurface=0.9
Diffuse
Specular

## Slide 23: Dimensions and specifications of test apparatus
Major specifications
Ultimate vacuum | 1.8×10-6Pa
2nd shield temperature | 3.3 – 4.1K
Required time to evacuation and cooling down | 24 hours
Required time to warm-up and repressurization | 13 hours
*The values above is for blank test.

## Slide 24: Heat leak test results
Ratio of heat leak to radiation heat transfer rate of a sample (Qleak/Qrad,s)
In the case of es=0.8:
Qleak/Qrad,s = 4-6% at maximum
In the case of es=0.3:
Qleak/Qrad,s = 11-16% at maximum
Heat leak test results

## Slide 25: Test sample set up