## Slide 1: Long-term Operation of the Superconducting Triplet Quadrupoles (STQ) with Cryocoolers for BigRIPS In-flight Separator at RIKEN
Introduction
BigRIPS & BI-beam Delivery line
STQ with cryocoolers
Operational Experiences
  Maintenance
  Long-term Operation of Cryocoolers
  Troubles
KUSAKA Kensuke
RIKEN Nishina Center, BigRIPS Team
Contents

## Slide 2
Introduction
Standard Operation of
SC Magnets
Thermal Cycle Operation
Operation of SC Magnets with Cryocoolers
Excitation
Fabrication
Cold Maintenance
Excitation
Cold Maintenance
Repeat . . . .
Never warm up SC magnet
Continuous operation of cryocoolers
longer than 10 years

## Slide 3: BigRIPS and RI-beam Delivery Lines
Unique large aperture (ID 240mm) 100m-long Beam Lines
STQ18 on circular motion stage
No Transfer Line needed
Flexible in beam-line design
STQ18 is used in 2 beam lines
STQ with Cryocoolers
consists of 22 “Stand-alone”
Superconducting Triplet Quadrapoles

## Slide 4: STQ Magnet Specification
Conservative LTS Quads
Q500
Q500
Q800/Q1000
Superferric
Total Cold mass ~ 9.4/10.2 tons
Triplet in single cryostat
Flat Racetrack Coil
NbTi
LHe bath cooling
 | Q500/Q800/Q100
Effective length [m] | ~ 0.54 / ~0.84 / ~1.04
Pole-tip dia. [mm] | 340
Warm bore dia. [mm] | 240
Length of iron [mm] | 440 / 740 / 940
OD of yoke [mm] | 960
 | Q500/Q800/Q100
Field gradient [T/m] | 14.1
Conductor size [mm] | f1.15
Cu/super ratio | 6.5
# of turns | 1405
Ic/Iop | 3.5
Nominal current [A] | 142 / 137 / 135
Current density [A/ mm2] | 115 / 110 / 109
Stored energy [MJ] | 0.13 / 0.19 / 0.27

## Slide 5: STQ with Cryocoolers
HTSC PLs
4K Cooler
Shield & PL Cooler
GM/JT cooler
2.5W@4.3K
Single-stage GM cooler
90W@80K
Small Heat Load < 2W
Large Cold Mass ~10 ton
SHI CG308
SHI CG110
He Vessel
LHe ~500 L
Cryocooler with Gas Driven Displacer
Key of Cold Maintenance
room-temp. displacer can be put
in cold cylinder
Unique Thermal Cycle operation

## Slide 6: Precooling
Room T (~300K) to 80 K
80 K to 4.5 K
GM/JT Cooler
Mounting
after precooling
2005 ~ 2006
Continuous
Operation
of
Cryocoolers
Transferring liquid N
from CE tank
3000 ~ 4000 L
Transferring liquid He
from Dewar
~ 2000 L

## Slide 7: Operation of Cryocoolers & STQ
Pressure of He vessel
Kept in 7 ~8 kPaG
> P atmospheric
Heater in vessel（2W) ON/OFF control
P
Back-Up of Electric Power & Cooling Water
Heater Output
Regular maintenance of Cryocoolers (yearly)
2007 ~ Present
Cooling capacity of GM/JT cooler (2.5W)
Fine tuning by JT valve (once ~ twice /year)
6500 kW- Power Class
Gas Turbine Cogeneration
Electric Power -> RIBF CGS & TEPCO
2 independent Cooling Water Systems
Spare Cryocoolers
& Compressors
Excitation of Magnets
PL temp. interlock < 70K
LHe level interlock > 80%
He loss occurs when P > 20 kPaG
Cracking Pressure of safety check valve

## Slide 8: Cryocooler Maintenance
Displacer Replacement
JT-circuit flushing
GM Cooler Maintenance（Cold Maintenance）
helium-purged gas bag
Exchange displacers
GM/JT Cooler Maintenance
Till 2008 warm maintenance
Magnets in Cold State
Well-established by
Cryocooler maintainers
From 2009 to present
Heater-bar maintenance
2. JT-flushing with heater
in Exp. displacer cylinder
without dismount GM/JT cooler from cryostat
1. Take out cold displacer
3. Insert new displacer
70~80 L of He loss
less than 10 L of He loss

## Slide 9: A Long-term Trend of Temperatures, LHe levels, etc.
e.g. STQ7 (from Jan. 2008 to Jul. 2018)
PL & GM head Temperature
Scheduled (yearly) Maintenances
by blocked gas flow in displacer
Constant LHe Level & well-cooled (< 70K) PL Temp.
Safe excitation of Magnets
Discharge & Suction Pressures of GM comp.
GM cooler degradation
LHe Level, Shield Temp., excess cooling capacity (ave. heater output P)
Rise of GM head, PLs & Shield Temp.

## Slide 10
e.g. STQ17 (from Jan. 2008 to Jul. 2018)
Unexpected rapid rises of GM head Temp.
Scheduled (yearly) Maintenances
flushing comp.
Rapid Increase
Degradation of cooling capacity of GM/JT cooler
GM/JT cooler replaced
JT-circuit blockage
T4K ~ 4.75K (4.3~4.5K)
Temp. of Expander Head
Texp ~ 16 K (<14K)
Unscheduled maintenances
Cooling capacity recovered after warm-up
He loss, pvessel > 20 kpaG
A Long-term Trend of Temperatures, LHe levels, etc.

## Slide 11: Unscheduled Maintenance of GM coolers & replaced GM/JT coolers
Year | STQ system
2008 | 
2009 | STQ8
2010 | STQ8, STQ11
2011 | STQ9
2012 | STQ6, STQ8, STQ9, STQ14, STQ18
2013 | STQ16
2014 | 
2015 | STQ17
2016 | STQ9, STQ10, STQ25
2017 | STQ21
operation time
from last maintenance
1,200~6,700 h
Unscheduled Maintenances
of GM coolers
 | Date of Stop | Operation time from maintenance (total) | Reason
STQ17 | May 2011 | 6,476 (38,773) h | TJT ~ 4.8K
STQ17 | Jun 2012 | 6,719 (47,913) h | Texp ~ 16K
STQ13 | Mar 2013 | 5,113 (69,439) h | Texp ~ 23K
STQ15 | Nov 2014 | 1,518 (82,003) h | Texp ~ 18K
STQ10 | Nov 2014 | 1,855 (83,017) h | Texp ~ 17K
STQ11 | Jun 2017 | 5,983 (106,985) h | TJT ~ 4.6K
TJT 4.3 ~ 4.5 K
Texp < 14K
Normal temp. of GM/JT cooler
Few hours beam time stops
Replaced (deteriorated) GM/JT coolers

## Slide 12: Hardware failures
1st Stage
Expander
2nd Stage
JT valve
Adsorber
Re-condensing
Exchanger
JT Tuner
(Micrometer)
JT EX (1st)
JT EX (2nd)
JT EX (3rd )
ADS
ADS
Ballast Tank
Comp
Comp
Comp
P-cntrl
valve
P-cntrl
valve
Surge Tank
check
valve
Oil
SP
Oil
SP
GM compressor (STQ14) top 60,086 h
Schematic Diagram of GM/JT cooler system
Gass flow decreased 20%, (Oil SP, P-control valve )
GM/JT compressor (STQ24) top 17,245 h
Malfunction of Exp. Rotary Comp. (foreign substance)
Whole Unit
Replaced
Exp. Comp
Replaced
GM/JT compressor (STQ11) top 71,775 h
Many sudden Stops (PLC error)
Aging deterioration of sensors
(P-sensor in HP line)
sudden stop of comp.
GM/JT comp.
GM comp.
STQ16 top 70,087 h
STQ11 top 82,448 h
STQ20 top 100,023 h
STQ21 top 100,563 h
STQ23 top 89,603 h
STQ6 top 103,855 h
PLC Replaced
JT valve of GM/JT cooler
STQ18 top 69,190 h
STQ19 top 78,737 h
cold head T didn’t change
JT valve Replaced

## Slide 13: Summary & Conclusion
We have developed STQ with Cryocoolers
We improved the maintenance of GM/JT coolers
  Heater bar maintenance to reduces He loss
We operate 22 STQ systems for more than 10 years without warming up magnets.
Yearly maintenance is essential to avoid degradation of cooling capacity.
Unexpected rise of GM head temperature is under investigation. (impurity of coolant gas He)
Hardware failures by aging are becoming noticeable.
Preparing spare coolers & compressors are essential.
Solid infrastructure (electric power & cooling water) is important for long-term operation.

## Slide 14: RIBF Cryogenics
SRC
IRC
SAMURAI
100m-long beam lines formed by
Stand-alone Superconducting Quads
5 STQ cooled by
Cryogenic Plant
BigRIPS
1st Stage
Linde TCF200
Linde TCF50S
Charactorized by
Beam Heat load
to Cryogenic System
22 STQ’s
Cooled by Cryocoolers
SHI CG308
Large Heat Load Fluctuation!
3T 880mm-gap
Superconducting Dipole
Cooled by Cryocoolers

## Slide 15: He Loss during maintenance
1.5day
Blow to reduce P of He vessel
Stop of compressor longer than 30h
Compressor Stop
& Dismount cryocooler
Mount &
Restart compressor
P [kPaG]
Check valve
P Relief =3psi
LHe Loss
Level 82% -> 59%
~60L
70~100L of LHe transfer / 1 STQ
24h warm-up for JT flushing
STQ14 4K ref.

## Slide 16: Improvement of GM/JT Cooler Maintenance II
He vessel Pressure during maintenance
Level 87% -> 85%
LHe Loss
~ Few L
50~70L of LHe transfer / STQ in 2~3 years
STQ17
Heater-bar maintenance
JT flushing without dismounting
cryocooler from cryostat
Warm-up JT circuit by heater-bar
in Expander Cylinder
~5h
Compressor
Stop
Restart
P [kPaG]
Save ~ 1000L of LHe / year !