## Slide 1: Conceptual design of the Cryogenic System of Comprehensive Research Facility for Key Fusion Reactor Core Systems
Liangbing Hu
huliangbing@ipp.ac.cn
Sep.4. 2018

## Slide 2: Outline
Background
Cryogenic facility
Conclusion and Perspective
Cryogenic supported system
Research Platform of Key technology for helium cryogenic

## Slide 3: Key components: magnets and divertor
The facility encompasses two synergistic research systems for :
1. Superconductor and magnets capable of max. field of 16.7T
2. Divertor physics and engineering capable of a max.plasma flux of 1024/m2s.
CFETR
Magnets
Divertor
“Comprehensive Research Facility for Key Fusion Reactor Core System” in National Grand Project by Chinese government will be build to ensure their overall performance, reliability and safety in CFETR.
To break through the bottlenecks of magnets and divertor for experimental research and engineering development.
( Chinese Fusion Engineering Test Reactor)

## Slide 4: Configuration of magnets system
Magnets system
Magnet research platform
Cryogenic research platform
Material research platform
Conductor research platform
Power research platform
Parameters | 
Magnet dimension: | 10m x 15m
Max field | 16.7T
Cool Capacity | 7.2kW@4.5K
Max current | 120kA
Research of superconductor and high magnet characteristic, ensure the safety and reliability of superconducting magnets for fusion reactor

## Slide 5: Scope of cryogenic system
Superconducting magnets test
Cryogenic research platform
Cryogenic public service platform
Superconductor performance
material performance
Divertor system
NBI system
ICRH system
Storage system
Compressor station
Cold box
Distribution system
Item | parameters
Capacity | 7.2kW@4.5K
Temperature | 1.8K，3K，4.5K
Pressure | 1600Pa~2MPa
Mass flow rate | >500g/s
Superconducting magnets test device
Divertor test device
Cryogenic research platform
Cryogenic public service platform
Cryogenic users

## Slide 6: Cryogenic system

## Slide 7: 3kW/4.5K refrigerator
The refrigerator is based on Claude cycle with LN2 precooling and two turbine in series. The 3K stage is obtained by two cold compressors in series.
The Capacity is 3kW/4.5K（1kW/3K） . FOM is 20%.
TYPE | refrigerating mode(4K) | refrigerating mode(3K) | 50K-80K thermal shield | Mass flow rate of SHe
HRL 3000 | 3000W/4.5K | 1.2kW/4.5K+1.0kW/3K | 15.5kW/50K | 200g/s

## Slide 8: 1kW/4.5K refrigerator
The refrigerator is based on Claude cycle with LN2 precooling and two turbine in series.
The Capacity is 1050W/4.5K on pure refrigeration mode. Liquid rate is 300L/hr, FOM is 21.5%.
TYPE | LN2 precool |  | No LN2 precool | 
HRL 1000 | Capacity | Liquid rate | Capacity | Liquid rate
 | 1000W/4.5K | 300L/h | 600W/4.5K | 150L/h

## Slide 9: 200W/4.5K refrigerator
The refrigerator is based on Claude cycle with LN2 precooling and two turbine in series.
The Capacity is 210W/4.5K on pure refrigeration mode. Liquid rate is 60L/hr, FOM is 12.6%.
TYPE | LN2 precool |  | No LN2 precool | 
HRL 200 | Capacity | Liquid rate | Capacity | Liquid rate
 | 200W/4.5K | 60L/h | 100W/4.5K | 30L/h

## Slide 10: 250W/1.8K refrigerator
1.8K cycle is obtained by cold compressor and warm vacuum compressor mixed mode.
It will provide the 110K, 80K, 20K, 4.5K interface for the research platform of heat mass and transfer in cryogenic.
ITEM | Temperature | Capacity | Pressure | Mass flow rate
Parameters | 1.8K~300K | 250W/1.8K 500W/2K | 1600Pa~2MPa | 10~25g/s

## Slide 11: Cryogenic research platform
Major function:
Develop and make research on cryogenic system for CFETR and future fusion reactor and promote its application in industry.
Research of heat and mass transfer in cryogenic
Key technology of helium cryogenic
Research of fluid machinery properties in cryogenic
Simulation of cryogenic process, control and safety
Test of cryogenic measurement and calibration

## Slide 12: Heat and mass transfer in cryogenic
He II & He I Phase change mechanism
Cryogenic fluid flow and heat transfer
Cryogenic heat exchanger
Oil removal system experiment facility
Separation of impurities gas from He
Cryogenic flow & heat transfer research platform
Mechanism study of He II phase change heat transfer in superconducting magnets and transient thermal response under pulsed heat load
Cryogenic fluid and heat transfer
He II&He I Phase change mechanism
Cryogenic heat exchanger
Oil removal system experiment facility
Separation of impurities gas from He

## Slide 13: Fluid machinery properties in cryogenic
Turbine expander
Cold compressor
Helium circulator
4.5K helium refrigerator
1.8 K refrigeration
unit
Test cryostat for Cold compressor
Test cryostat for
subcool heat exchanger
The cryogenic fluid machinery properties in cryogenic will be to develop the large-flow turbine expander, cold compressor and helium circulator, and realize the in-house production of key components in large scale helium cryogenic system.

## Slide 14: Simulation of cryogenic process, control and safety
Refrigeration Process Design
Thermodynamic calculation, Equipment parameters
Control Strategies Design
Control sequences , Controller Parameter Tuning,
Virtual commissioning
Performance Optimization
Operational optimization in Off-design condition
Management under abnormal conditions
Fault diagnosis, Smooth for heat-load
System Commissioning
Operator training, Refrigeration cycle re- optimization
It will be an efficient platform to realize rapid design and stable operation of cryogenic system, and have guiding significance during the life-cycle of cryogenic system .
Dynamic simulation of transient characteristics of pulse heat load on cryogenic system and collaborative control to smooth the influence of the load.

## Slide 15: Test of cryogenic measurement and calibration
Cryogenic measuring and calibration experiment device under magnet field
Cryogenic performance test device for sealing material and structure
Cryogenic flow meter test device under magnet field
The service platform is established for calibration and test of cryogenic precision measurement instruments, high-voltage insulation performance, insulating materials under 1.8K, and cryogenic performance test of sealing materials and sealing structures, flow characteristics, heat leakage characteristics of cryogenic valve.

## Slide 16: Time table
Project | 2018 |  |  |  | 2019 |  |  |  | 2020 |  |  |  | 2021 |  |  |  | 2022 |  |  | 
 | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 4 | 1 | 2 | 3 | 4
Cryogenic system |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Engineering design of helium refrigeration |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Manufacture of cold box and distribution valve box |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Assemble of helium refrigerator |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Commission and assemble of helium refrigerator |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Commission and system integration of cryogenic system with load |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Engineering design of cryogenic research platform |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Manufacture of sub system |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Assemble of sub system |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Commission and test of sub system |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Optimization and acceptance of project |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 

## Slide 17: Location of the facility
Science island
Hefei Comprehensive National Science Center, Anhui ,China
16,500 m2
Cryogenic system
New Facility

## Slide 18: Conclusion and perspective
Conceptual design of the project have been finished and the proposal has been submitted to be review.
The preliminary engineering design will be evaluated.
A large-scale Cryogenic Technology and Engineering center will be built in 5 years which can provide the cryogenic technology and test for fusion reactor and other users.
Collaboration and technical supports will be welcome.

## Slide 19
Thanks for your attention！